## software resources
Useful links to download resources
k0-INRIM software | https://github.com/marcodiluzio/k0-INRIM/releases/tag/v2.1
k0-INRIM user manual | https://github.com/marcodiluzio/k0-INRIM_documentation/blob/4f16d0917c2f4388aec94dd1798a87ade837c7e5/k0-INRIM_user_manual.pdf
k0-INRIM getting started | https://github.com/marcodiluzio/k0-INRIM_Getting-started

## Settings
Set the software settings as reported here
Location: Utility -> Settings (hover buttons with the mouse and see at the bottom left of the current window to get hints of locations)
k0 database | k0-database_2020_08_13_m.xls
ΔE / keV | 0.35
items per page | irrelevant (depends on your screen size)
max uncertainty of calibration parameters | 40%
statistical limit for calibration spectra | 5%
statistical limit for standard spectra | 15%
statistical limit for sample spectra | 40%
non-certified target uncertainty | 20%
default uncertainty for tc and tl / s | 0.2
default uncertainty for td / s | 2
always look for spectrum file | True
Press the 'apply changes and restart' button

## Detector characterization I
Create a new gamma source with the following data and save it
Location: Utility -> Browse database -> gamma source database
Add a new source by pressing the 'add a new source' button
source name |  | irrelevant
certificate date |  | 21/09/2021 17:45:14
Add the following information with the 'emitter' drop-down menu, 'activity / Bq' and 'half-life / s' entry fields and 'add or update an emitter' button in the bottom-left region of the window
emitter | activity / Bq | half-life / s
Ba-133 | 298769.7 | 332959680 | once the emitter is added, delete the emissions at 30.6, 31.0, 35.0 and 35.9 keV through the drop-down menu and the 'delete emission' button in the bottom-right region of the window
Am-241 | 15493 | 13651536960 | once the emitter is added, delete the emission at 26.3 keV through the drop-down menu and the 'delete emission' button in the bottom-right region of the window
Cd-109 | 33751.4 | 39864960 | once the emitter is added, delete the emissions at 25.0 and 25.5 keV through the drop-down menu and the 'delete emission' button in the bottom-right region of the window
Eu-152 | 292801.3 | 426556800 | once the emitter is added, delete the emissions at 39.5, 40.1, 45.5 and 46.7 keV through the drop-down menu and the 'delete emission' button in the bottom-right region of the window
Co-57 | 9174.6 | 23474880
Cr-51 | 9451.2 | 2393280
Au-198 | 13790.6 | 233280 | once the emitter is added, delete the emissions at 68.9, 70.8 and 675.9 keV through the drop-down menu and the 'delete emission' button in the bottom-right region of the window
Cs-137 | 16011.4 | 949233600 | once the emitter is added, delete the emissions at 31.8, 32.2 and 36.5 keV through the drop-down menu and the 'delete emission' button in the bottom-right region of the window
Zn-65 | 5281.2 | 21072960 | once the emitter is added, delete the emission at 511.0 keV through the drop-down menu and the 'delete emission' button in the bottom-right region of the window
Check if the following 7 g-emissions are marked as COIfree True (the box labeled COIfree in the bottom-right region of the window should be checked)
Am-241, 59.54 keV
Cd-109, 88.03 keV
Co-57, 122.06 keV
Cr-51, 320.08 keV
Au-198, 411.8 keV
Cs-137, 661.66 keV
Zn-65, 1115.64 keV
Press the 'confirm changes to selected source' button

## Detector characterization II
Create a new detector characterization and save it
Location: Characterization -> New detector characterization
name | irrelevant (any name without special symbols is valid)
detector | irrelevant (any name without spaces and special symbols is valid)
source | recall the gamma source previously saved, open the selection window through the 'select emissions' button and deselect the emissions at 563 keV and 1005 keV
Define counting positions by setting the distances and recalling corresponding spectra
positions | reference
d / mm | 203.4
spectra (recalled through the button 'Add spectra to the current counting position')
200_Am.csv
200_Au.csv
200_Ba.csv
200_Cd.csv
200_Co.csv | open the peaklist through the 'show the peaklist of the selected spectrum' button and select the Co emission at the 122.05 keV line
200_Cr.csv
200_Cs.csv
200_Eu.csv | open the peaklist through the 'show the peaklist of the selected spectrum' button and select the Eu emission at the 121.77 keV line
200_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 183.4
spectra (recalled through the button 'Add spectra to the current counting position')
180_Am.csv
180_Au.csv
180_Cd.csv
180_Co.csv
180_Cr.csv
180_Cs.csv
180_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 163.4
spectra (recalled through the button 'Add spectra to the current counting position')
160_Am.csv
160_Au.csv
160_Cd.csv
160_Co.csv
160_Cr.csv
160_Cs.csv
160_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 143.4
spectra (recalled through the button 'Add spectra to the current counting position')
140_Am.csv
140_Au.csv
140_Cd.csv
140_Co.csv
140_Cr.csv
140_Cs.csv
140_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 123.4
spectra (recalled through the button 'Add spectra to the current counting position')
120_Am.csv
120_Au.csv
120_Cd.csv
120_Co.csv
120_Cr.csv
120_Cs.csv
120_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 103.4
spectra (recalled through the button 'Add spectra to the current counting position')
100_Am.csv
100_Au.csv
100_Cd.csv
100_Co.csv
100_Cr.csv
100_Cs.csv
100_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 83.4
spectra (recalled through the button 'Add spectra to the current counting position')
80_Am.csv
80_Au.csv
80_Cd.csv
80_Co.csv
80_Cr.csv
80_Cs.csv
80_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 63.4
spectra (recalled through the button 'Add spectra to the current counting position')
60_Am.csv
60_Au.csv
60_Cd.csv
60_Co.csv
60_Cr.csv
60_Cs.csv
60_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 43.4
spectra (recalled through the button 'Add spectra to the current counting position')
40_Am.csv
40_Au.csv
40_Cd.csv
40_Co.csv
40_Cr.csv
40_Cs.csv
40_Zn.csv
positions | add counting position through the button 'Add a counting position' (the name is irrelevant but can be changed through the 'rename a counting position' button)
d / mm | 23.4
spectra (recalled through the button 'Add spectra to the current counting position')
20_Am.csv
20_Au.csv
20_Cd.csv
20_Co.csv
20_Cr.csv
20_Cs.csv
20_Zn.csv
Calculate peak-to-total
Select the counting positions at 83.4 mm distance and click on the 'Evaluate PT for the current position' button
background | recall the spectrum Background.csv
junction E / keV | 250
click 'start elaboration' and 'confirm elaboration' buttons to elaborate and save the peak-to-total results
click 'elaborate detector characterization' and 'save the current detector characterization' to finalize and save the characterization

## Preparation of samples
Create and save two material files: one for the standard (monitor) and the other for the Reference Material (RM)
Location: Utility -> Browse database -> material database
standard (monitor)
Add a new material by pressing the 'add a new material' button
material name | irrelevant (any name without special symbols is valid)
description | irrelevant
material type | irrelevant
physical state | irrelevant
Composition (add elements by selecting them from the drop-down menu and introducing mass fraction values in entry fields at the bottom of the window)
element | w / g g-1 | u(w) / g g-1
Au | 0.0009737 | 2.43e-06
Click 'update changes to this material' to save the modifications
reference material
Add a new material by pressing the 'add a new material' button
material name | irrelevant (any name without special symbols is valid)
description | irrelevant
material type | Reference Material
physical state | irrelevant
Composition (add elements by selecting them from the drop-down menu and introducing mass fraction values in entry fields at the bottom of the window)
element | w / g g-1 | u(w) / g g-1
Ca | 0.00196 | 9e-05 | 4.592%
Mg | 0.00118 | 8e-05 | 6.78%
K | 0.0561 | 0.0013 | 2.317%
Na | 0.0265 | 0.00067 | 2.528%
As | 5.7e-05 | 1.6e-06 | 2.807%
Fe | 0.000539 | 7.5e-06 | 1.391%
Sr | 0.0121 | 0.00029 | 2.397%
Zn | 0.00157 | 5.8e-05 | 3.694%
Click 'update changes to this material' to save the modifications

## Neutron irradiation I
Create a new irradiation facility characterization and save it
Location: Characterization -> Flux evaluation
channel |  |  | irrelevant (any name without spaces and special symbols is valid)
pos |  |  | irrelevant (any name without spaces and special symbols is valid)
end of irradiation date |  |  | 2021-07-06 15:40
ti / s |  |  | 18000
calibration |  |  | recall the one previously saved
Add information for neutron flux monitors
emitter | mass / g | u(mass) / g | Gs / 1 | u(Gs) / 1 | Ge / 1 | u(Ge) / 1 | condition
Au-198 411.8 | 1.506e-05 | 1.95e-08 | 1 | 0 | 1 | 0 | lowest unc.
Nb-97m 743.4 | 0.01709 | 1e-05 | 1 | 0 | 0.982 | 0.003 | earliest
Zr-95 756.7 | 0.01709 | 1e-05 | 1 | 0 | 0.996 | 0.003 | lowest unc.
set triple monitor counting distance |  |  | 203.4
spectra (recalled through the button 'Add spectra for flux evaluation')
flux_monitor1_200_08072021.csv
flux_monitor1_200_19072021.csv
Click 'compute flux parameters' and 'confirm obtained values' to elaborate and save the results

## Neutron irradiation II
Create a new neutron flux gradient characterization and save it
Location: Characterization -> Gradient evaluation
channel |  |  | recall the one previously saved
end of irradiation date |  |  | 2021-07-06 15:40
calibration |  |  | recall the detector characterization previously saved
monitor |  |  | Au-198 411.8
Add information for flux gradient monitors
| spectrum | mass / g | u(mass) / g | Gs / 1 | u(Gs) / 1 | Ge / 1 | u(Ge) / 1 | u(Δd) / mm | ΔL / mm | u(ΔL) / mm
↑ | flux_monitor1_200_08072021.csv | 0.01547 | 2e-05 | 1 | 0 | 1 | 0 | 0.2 | 10 | 0.2
- | flux_monitor0_200_14072021.csv | 0.01581 | 2e-05 | 1 | 0 | 1 | 0 | 0.2
set monitor counting distance |  |  | 203.4
set position name (within the irradiation channel) |  |  | irrelevant (any name without spaces and special symbols is valid)
Click 'evaluate' and 'save elaboration' to elaborate and save the flux gradient value

## Data elaboration
Perform the analysis elaboration to get mass fraction results
Location: Characterization | recall, from the corresponding drop-down menu, the detector characterization previously saved
Location: Analysis name | irrelevant (any name without special symbols is valid)
Location: Irradiation -> New irradiation
irradiation code | irrelevant (any name without special symbols is valid)
channel name | recall the previously saved flux parameters through the 'select flux parameters from database' button
position name | recall the previously saved beta parameter through the 'select beta parameter from database' button
end of irradiation date | 2021-07-06 15:40
irradiation time / s | 18000 | 17
Click 'confirm irradiation data' to save the irradiation data
Location: Background
recall background spectrum | background.csv
Location: Standard
recall standard spectrum | standard_200_14072021.csv
open the peaklist window through the 'standard peaklist' button, select the radiobutton k0 in the line at 411.7 keV and select the string 'Au-197 411.8 keV' from the column emitter to adopt it as a monitor
open the define material window through the standard 'define material' button to add information for standard
material | select the material previously saved for the monitor (the one with the Au in it)
| value | uncertainty
weighted mass / g | 0.01581 | 2e-05
moisture / % | 0 | 0
Gth / 1 | 1 | 0
Gepi / 1 | 1 | 0
sample height / mm | 0.1 | 0
density / g cm-3 | 1e-06 | 0
click the button at the bottom-left of the window to confirm the settings
Set counting position
horizontal slider | 203.4
Δd / mm | 0.4
u(Δd) / mm | 0.14
Location: Sample
recall sample spectra | Analyte_sample_80_10072021.csv
CRM_80_09072021.csv
select the Analyte_sample_80_10072021 from the sample drop-down menu
open the define material window through the sample 'define material' button to add information for sample
material | leave the drop-down menu unselected
| value | uncertainty
weighted mass / g | 0.20165 | 2e-05
moisture / % | 3 | 0.7
Gth / 1 | 1 | 0
Δl / mm | 8 | 0.2
sample height / mm | 1.8 | 0.1
density / g cm-3 | 1.427 | 0.071
check the box with label 'find peaks' for automatic peak identification
click the button at the bottom-left of the window to confirm the settings
select the CRM_80_09072021 from the sample drop-down menu
open the define material window through the sample 'define material' button to add information for sample
material | select the material previously saved for the CRM (the one labeled as Reference Material)
| value | uncertainty
weighted mass / g | 0.168 | 2e-05
moisture / % | 0 | 0
Gth / 1 | 1 | 0
Δl / mm | 5 | 0.2
sample height / mm | 1.5 | 0.1
density / g cm-3 | 1.427 | 0.07
check the box with label 'find peaks' for automatic peak identification
click the button at the bottom-left of the window to confirm the settings
Set counting position
horizontal slider | 83.4
Δd / mm | 0.4
u(Δd) / mm | 0.14
Location: Results
Click the 'Elaborate' button to perform the analysis
Create a Microsoft Excel worksheet through the 'process excel output' button with 'group by spectra' and 'sort by Z number' options checked to save the results
Have a look to the other two methods of visualizing and checking the quality of results through 'overview' and 'perform a validation inspection' buttons, respectively

## expected results
Here are reported the expected results obtained from the analysis
Analyte sample
target | emitter | E / keV | wa / g g-1 | u(wa) / g g-1
Na | Na-24 | 1368.6 | 0.033055 | 0.00083302
K | K-42 | 312.7 | 0.059412 | 0.0059882
K | K-42 | 1524.7 | 0.059098 | 0.0014891
Co | Co-60 | 1173.2 | 6.982e-06 | 1.9044e-06
Cu | Cu-64 | 511.0 | 0.0019235 | 6.1943e-05
Zn | Zn-69m | 438.6 | 0.00048574 | 0.00010392
As | As-76 | 657.1 | 1.4699e-05 | 2.0793e-06
Br | Br-82 | 554.3 | 9.7067e-05 | 2.8198e-06
Br | Br-82 | 698.4 | 9.4703e-05 | 3.0468e-06
Br | Br-82 | 776.5 | 9.5684e-05 | 2.7006e-06
Br | Br-82 | 827.8 | 9.8492e-05 | 3.2277e-06
Br | Br-82 | 1044.0 | 9.0242e-05 | 2.6763e-06
Br | Br-82 | 1317.5 | 9.8268e-05 | 3.1311e-06
Br | Br-82 | 1474.9 | 9.3327e-05 | 2.8895e-06
Sb | Sb-122 | 564.2 | 2.0014e-06 | 7.9296e-08
La | La-140 | 1596.2 | 1.9091e-07 | 4.8932e-08
Au | Au-198 | 411.8 | 1.1358e-06 | 1.2542e-08
Th | Pa-233 | 398.5 | 2.4677e-05 | 8.8841e-06
CRM
target | emitter | E / keV | wa / g g-1 | u(wa) / g g-1
Na | Na-24 | 1368.6 | 0.026934 | 0.00064879
K | K-42 | 312.7 | 0.05702 | 0.0023199
K | K-42 | 1524.7 | 0.065794 | 0.0015323
Zn | Zn-69m | 438.6 | 0.0063768 | 0.00020514
Zn | Zn-65 | 1115.5 | 0.0015952 | 7.7277e-05
As | As-76 | 559.1 | 5.4663e-05 | 4.9983e-06
As | As-76 | 563.2 | 4.3043e-05 | 5.0093e-06
As | As-76 | 657.1 | 5.5731e-05 | 5.3001e-06
As | As-76 | 1212.9 | 5.9613e-05 | 6.8909e-06
As | As-76 | 1216.1 | 5.1783e-05 | 5.065e-06
Sr | Sr-85 | 514.0 | 0.0090584 | 0.0027247